Operating Documentation Direction 5193755-800, Rev# 11 BrightSpeed (Ltd. Access) Service Methods - Adv Module Version 1.0, Version Date 2/16/07

# No Load kV Test (LS16, Ultra, Plus)

(Click here for Summary Version)

## 1  Personnel Requirements

| Required Persons | Preliminary Reqs | Procedure | Finalization |
| --- | --- | --- | --- |
| 1 |  | 30-60 depending on outcome | Included |

## 2  Overview

Objective

- The purpose of this test is to verify that the HV power inverter and HV tank are working properly. The exposure is taken as in application mode except that no filament drive nor anode rotation is running. Verification is made that the inverter currents are correctly set and that kV regulation is operating properly. As no filament drive is applied, no x-rays are generated.

This test also allows to separate generator from HV cable or x-ray tube problem by running it with or without HV cables plugged on the HV tank ( see Notice) This test requires manual operation.

Summary

- Remove HV Candlesticks from tank receptacles

- Fill tank receptacles with oil

- Run diagnostic

- Clean up

| Last Revised: | September 2003 |
| --- | --- |

## 3  Preliminary Requirements

### 3.1  Consumables

| Item | Qty | Effectivity | Part# | Manufacturer |
| --- | --- | --- | --- | --- |
| Tube Oil | 40 mL (1.4 oz) may be necessary | - | - | - |

### 3.2  Safety

| NOTE: | See the WARNING and NOTICE statements in Section 4. |
| --- | --- |

### 3.3  Required Conditions

| Condition | Reference | Effectivity |
| --- | --- | --- |
| Generator input line in an acceptable range (670VDC +/-10%). | - | - |
| PPC kV control board alive and running : DS17 - DS24 Leds are lit successively or a combination of them blink. | - | - |
| inverter gate_cmd board DC supply present : DS300 neon is lit. | - | - |
| inverter quad snubber board DC supply present : DS1-DS2 leds are lit. | - | - |
| Shorted inverter test passed without failure. | - | - |

## 4  Procedure

| NOTE: | First execute the test with the HV cables connected. |
| --- | --- |

- Launch the [GENERATOR TOOL].

- Launch the [NO LOAD KV DIAGNOSTIC] tool.

- Run the diagnostic. Verify:

  - No error is reported on the console

  - inverter gate_cmd board Leds DS101, DS 102, DS201, DS202 are lit: IGBTs gate drive supply is working properly.

If this test does not pass, then execute the test with the HV cables disconnected.

|  | WARNING |
| --- | --- |
|  | POTENTIAL FOR SHOCK. |
|  | VOLTAGE MAY BE PRESENT. |
|  | USE PROPER LOCKOUT/TAGOUT PROCEDURES BEFORE REPLACING COMPONENTS. |

- If needed, remove gantry side and front covers.

- Remove the JEDI FAN Assy.

- Remove HV cable candlesticks from tank receptacles.

|  | NOTICE |
| --- | --- |
|  | Equipment damage possible. You must fill the receptacles with oil to avoid arcing. |

- Fill tank receptacles with oil.

- Power on the generator using 120VAC Enable and HVDC Enable switches on the service switch panel. Also press [ESTOP DRIVES RESET].

- Launch the [GENERATOR TOOL].

- Launch the [NO LOAD KV DIAGNOSTIC] tool.

- Run the diagnostic. Verify:

  - No error is reported on the console

  - inverter gate_cmd board Leds DS101, DS 102, DS201, DS202 are lit : IGBTs gate drive supply is working properly

## 5  Finalization

- If the latter part of the test passes, along with the “Shorted Inverter Test ”, the system more than likely has a tube or HV cable problem.