| ERROR CODE | MESSAGE | PROBLEM | ACTION |
| --- | --- | --- | --- |
| E007 | Filament Fault | Power supply has detected filament current <2 amps. | Check for open filament in X-ray tube. Check for poor connections in the cathode cable. Check fuses on filament board(s). |
| E008 | mA Fault | Power supply has detected a fault in the kV or mA output during an exposure and immediately terminated the exposure. This may be caused by arcing in the X- ray tube, arcing of the HV cables, or HT tank. | If arcing of the X- ray tube is suspected, check condition of tube. The X-ray tube may be damaged or simply require "seasoning". Refer to chapter 6 for tube seasoning procedure. If failure of HT tank is suspected, contact product support. |
| E009 | Power Supply Not Ready | Power supply is not ready to make an exposure. | Retry exposure. |
| E011 | MA During Exposure Too High | Generator CPU detected mA greater than allowed tolerance. | Recalibrate X-ray tube. |
| E012 | MA During Exposure Too Low | Generator CPU detected mA less than allowed tolerance. | Recalibrate X-ray tube. |
| E013 | Manually Terminated Exposure | Operator released exposure switch during exposure. | Re-take exposure if necessary. Check for faulty switch contacts or wiring. |
| E014 | AEC Back-up Timer - Exposure Terminated | AEC exposure exceeded allowed back up time. | Check exposure technique settings. Check that correct AEC chamber is energized. |

5.3.3 Error Messages (Cont)

| ERROR CODE | MESSAGE | PROBLEM | ACTION |
| --- | --- | --- | --- |
| E015 | AEC mAs Exceeded - Exposure Terminated | AEC exposure exceeded allowed back up mAs. | Check exposure technique settings. Check that correct AEC chamber is energized. |
| E016 | Tomo Back-up Timer - | Tomo exposure exceeded | 1. Check exposure |
|  | Exposure Terminated | back up time. | technique settings. 2. Increase tomo back up time if necessary. |
| E017 | Uncalibrated Exposure | Selected mA not calibrated | Recalibrate X-ray |
|  | Parameter | for selected kV. | tube. |
| E018 | Preparation Timeout | Generator has been in prep state too long. | Reduce length of time in prep state. |
| E019 | Anode Heat Limit | Selected parameters will cause X-ray tube to exceed its programmed anode heat limit. | Reduce parameters or wait for tube to cool. |
| E020 | Thermal Switch Interlock #1 | X-ray tube # 1 too hot and | Wait for X-ray tube # |
|  | Error | its thermal switch has opened. | 1 to cool. |
| E021 | Thermal Switch Interlock #2 | X-ray tube # 2 too hot and | Wait for X-ray tube # |
|  | Error | its thermal switch has opened. | 2 to cool. |
| E022 | Door Interlock | Door is open. | Close door. |
| E023 | Collimator Interlock Error | Collimator is not ready. | Check collimator. |
| E024 | Cassette Interlock Error | Cassette is not ready. | Check cassette. |
| E025 | Image Intensifier Safety Interlock Error | II safety is not ready. | Check II safety. |
| E026 | Spare Input Interlock Error | Spare input is not ready. | Check spare input. |
| E027 | Receptor Timeout Error - Receptor did not respond within the timeout period | The receptor is not responding | Wait for receptor to be ready, then retry exposure. |
| E028 | Prep Input Active During Initialization Phase | Prep input active during power on initialization phase. | Check prep switch and input for short circuit. |
| E029 | X-Ray Input Active During Initialization Phase | X-ray input active during power on initialization phase. | Check X-ray switch and input for short circuit. |
| E030 | Fluoro Input Active During | Fluoro input active during | Check fluoro switch |
|  | Initialization Phase | power on initialization phase. | and input for short circuit. |

5.3.3 Error Messages (Cont)

| ERROR CODE | MESSAGE | PROBLEM | ACTION |
| --- | --- | --- | --- |
| E031 | Remote Fluoro | Communication error | 1. Check remote |
|  | Communication Error | detected with remote fluoro control unit. | fluoro control cable for damage and proper connection. 2. Turn power off and then on to Reset Generator. |
| E032 | Console Communication | Generator has detected | 1. Check console |
|  | Error | error in communication to console. | cable for damage and proper connection. 2. Turn power off and then on to reset generator. |
| E033 | Warning Lithium Battery Voltage Low | Generator detects lithium battery voltage is low. | Replace lithium battery. |
| E034 | +12VDC Error | +12VDC rail is out of tolerance. | Check +12VDC rail. |
| E035 | -12VDC Error | -12VDC rail is out of tolerance. | Check -12VDC rail. |
| E036 | +15VDC Error | +15VDC rail is out of tolerance. | Check +15VDC rail. |
| E037 | -15VDC Error | -15VDC rail is out of tolerance. | Check -15VDC rail. |
| E038 | Calibration Data Corrupt | Generator detects corrupt | Re-calibrate X-ray |
|  | Error | calibration data. | tube(s). |
| E039 | AEC Data Corrupt | Generator detects corrupt AEC data. | Reprogram AEC data or set factory defaults. |
| E040 | Fluoro Data Corrupt | Generator detects corrupt Fluoro data. | Reprogram fluoro data or set factory defaults. |
| E041 | Receptor Data Corrupt | Generator detects corrupt receptor data. | Reprogram receptor data or set factory defaults. |
| E042 | Tube Data Corrupt | Generator detects corrupt tube data. | Reprogram tube data or set factory defaults. |
| E043 | High Voltage Error - kV | KV detected in non x-ray | Switch OFF |
|  | detected in non X-ray state | state. | generator. Prevent further use of generator. Call product support. |
| E044 | Invalid Communication Message | Received communication message not valid and ignored. | Reset error. |

5.3.3 Error Messages (Cont)

| ERROR CODE | MESSAGE | PROBLEM | ACTION |
| --- | --- | --- | --- |
| E045 | Communication Message Not Supported | Received message valid, but not supported by this system. | Reset error. |
| E046 | Mode Inhibited | Received message valid, but not allowed during present state. | Reset error. |
| E047 | Fluoro Timer Limit | Fluoro Timer has exceeded time limit. | Reset Fluoro timer. |
| E048 | Focus mismatch error | Focus selected does not match current focus enabled by power supply. | Check power supply interface cables between power supply and generator CPU board. |
| E049 | Not Enabled | Requested function not programmed to be enabled. | Reprogram to enable function. |
| E050 | Generator Limit Data Corrupt | Generator detects corrupt generator limit data. | Reprogram generator limit data or set factory defaults. |
| E051 | AEC Feedback Error (No Feedback Signal Detected) | Generator has detected no feedback signal from AEC device. | Check that X-ray tube is pointing at correct AEC device. Check AEC cable for damage and proper connection. |
| E052 | High Small focus filament current error in standby | Generator detects small focus filament current greater than limits in standby mode. | Check small focus filament board. |
| E053 | High Large focus filament current error in standby | Generator detects large focus filament current greater than limits in standby mode. | Check large focus filament board. |
| E054 | AEC Reference out of range | AEC reference has reached a maximum or minimum limit. | Re-adjust AEC calibration including density to operate within AEC range (O to 1OVDC). |
| E055 | No fields selected in AEC | AEC enabled but no fields are selected. | Select AEC field(s). |
| E056 | Receptor Disabled | All Receptors have no X-ray tube programmed. | Program receptor(s) with tube number. |

5.3.3 Error Messages (Cont)

| ERROR CODE | MESSAGE | PROBLEM | ACTION |
| --- | --- | --- | --- |
| E057 | AEC Stop Signal In Wrong State | AEC stop signal (P.T. stop signal) is active low indicating exposure is finished during prep state. | Check that P.T. ramp does not exceed P.T. reference during prep state. Check AEC device for proper operation. |
| E058 | Console Back-up Timer | Console has detected exposure exceeded backup time and terminated exposure. | Call product support. |
| E059 | Housing Heat Limit Exceeded | X-ray tube housing has exceeded limit. | Wait for tube to cool. |
| E060 | High kV Error | KV exceeds high KV tolerance level. | Check the output of the kV reference DAC on the generator CPU board. Measure the output of the generator with a dynalyzer or a non-invasive kVp meter. |
| E061 | Low kV Error | KV exceeds low KV tolerance level. | Check the output of the kV reference DAC on the generator CPU board. Measure the output of the generator with a dynalyzer or a non-invasive kVp meter. |
| E062 | EXP_SW Signal Active in Standby State | The EXP_SW signal on the generator Interface and generator CPU board is enabled when it should be disabled. | Call product support. |
| E063 | Factory Defaults Set | SW1 switch 8 on the generator CPU board is set to default the generator CPU NVRAM with factory defaults. | Set SW1 switch 8 to its non default position. The generator will not exit the initialization phase until this switch is set. |

5.3.3 Error Messages (Cont)

| ERROR CODE | MESSAGE | PROBLEM | ACTION |
| --- | --- | --- | --- |
| E066 | No Sync Signal | Pulsed fluoro has been requested, but no sync pulse is present. | Check that the imaging system is active, and cables are connected properly. Check JW22 on the generator interface board. |
| E067 | Power Supply Duty Cycle Limit | Power supply duty cycle limit exceeded. | Re-evaluate technique factors. Allow generator to cool if possible. If exposures are continued, serious generator damage may result due to overheating. |
| E070 | Generator software key error | Defective or missing GAL U29 on generator CPU board 734573. | Call product support for new GAL U29. |
| E071 | DAP overflow error | The accumulated DAP value exceeds the display limit. | Reset the DAP. |
| E072 | DAP device error | The DAP device is not functional. | Check the DAP wiring. Check the DAP interface board. |
| E073 | DAP data error | The DAP configuration data is corrupted. | Reset factory defaults. |
| E100 | Calibration Error - Maximum mA Exceeded | Maximum mA has been exceed during auto calibration. | Repeat auto calibration and/or decrease standby current. |
| E101 | Calibration Error - Calibration Data Table Exceeded | Auto calibration has exceeded data table length due to an excessive number of exposures. | Check to see if the filament standby current is too low. Retry auto calibration. |
| E102 | Calibration Error - Maximum Filament Current Exceeded | Maximum filament current for the selected focus has been exceeded. | Check to see if the maximum filament current limit can be increased. Retry auto calibration. |
| E103 | Calibration Error - Manually Terminated | Operator released exposure button during auto calibration. | Retry auto Calibration. |

5.3.3 Error Messages (Cont)

| ERROR CODE | MESSAGE | PROBLEM | ACTION |
| --- | --- | --- | --- |
| E104 | Calibration Error - No mA | No mA feedback detected during auto calibration. | Check power supply Interface cables between HF power supply and generator CPU board. |
| E105 | Calibration Error - Minimum mA Not Calibrated | Minimum generator mA was exceeded at start of calibration. This is usually caused by too high a filament standby current on the primary and or secondary filament. (Primary is the current filament being calibrated, secondary is the other filament and applies only to generators with two independent filament supplies). | Reduce filament standby current on primary and/or secondary filament. |