МАММО-ЗМТ TECHNICAL MANUAL

AUTOMATIC EXPOSUR

CONTROL

1	- THINGS TO KNOW BEFORE STARTING

- Mammography system

- Selecting film/screen combination

- Film Processor

2	- PROCEDURE PREPARATION

- Software installation

- Developer

- X-ray cassette for tests

- Film pack for tests

- Test phantom

- O.D. reference X-ray

- Time for carrying out procedure

- Anode filter combination

- Detector calibration

- Offset [o] Calibration

- Detector [r] Gain Calibration

- Connecting PC for calibration

- Selection of test layers

- Other Techniques and operating modes

- Automatic Molibdenum Rhodium filter

- Modification of the reference Density after calibration

- Maximum breast density limit

- Calibration concept

- Linearity check

- Unused Operating Technique and Film Screen

- Disabling specific Operating Technique and/or Film Screen

3	- CALIBRATION PROCEDURE

- Zero Point calibration

- Manual Density steps

- One point calibration

- Rhodium filter

- Maximum breast density limit

- AEC calibration files backup

MEDICAL SCIENTIFIC

Chapters Page21 of90

МАММО-ЭМ Г TLCIINICAL MANUAL

1 - THINGS TO KNOW BEFORE STARTING

The unit has an Automatic Exposure Control (AEC) that is among the most

advanced available on the market today.

Operation of the AEC is based on the effective breast density   measured by

means of a very short "pre-exposure"X-ray pulse.

This is the best way to operate the AEC exceeding performances of other

systems only based on breast thickness evaluation.

The radiation [r] read by the detector during pre-exposure represent breast

density and after being converted to a number that is inversely proportional to it,

will be used in the process of exposure control that will continue in a completely

numeric form, except for the Dose integration.

kV

The above figure shows a time schematization of the exposure concept in ZERO POINT technique, with completely automatic kV and mAs.

The parameter [p] derived directly from the [r], the density of the breast being

examined, is used as a pointer on two look-up tables containing the kV and

mAs necessary for obtaining an image with the right blackening and contrast to

fully satisfy any Quality Protocol in the Mammography field today.

A high value [p] corresponds to a LOW Density Breast and low values of kV

and mAs are delivered.

A low value [p] corresponds to a HIGH Density Breast and high values of kV

and mAs are delivered.

kV values can't be modified, but four different ranges are made available to the

operator, with different kV profiles suited to different requirements that go

beyond the possibilities offered by the well-known High Contrast (HC) and Low

Dose (LD) techniques.

From the control panel, the operator can not only select the technique that he

feels most suitable but he can also make fine-tuning in steps of 0.5kV on the kV

range selected.

MEDICAL SCIENTIFIC	Chapter 3 Page 22 of 90

МАММО-ЗМТ TECHNICAL MANUAL

The figure above is a diagram of the operating possibilities of the DEATA PLUS software with Mo target / Mo filter or Mo target with automatic Mo/Rh filters.

The large number of Film Screen and techniques available have to be intended to give the possibility to tune the system to the specific requirements of any single or multiple user and or different film that can be momentarily found in critical markets without loosing the possibility to go back to the previously used when newly available.

MEDICAL SCIENTIFIC

Chapter 3 Page 23 of 90

МАММО-ЗМТ TECHNICAL MANUAL

1.1	Mammography system

Mammography unit, Film, X-ray Cassettes and Screen, Film Processor and

Chemicals are the components that all together make up one SINGLE

SYSTEM, the final product of which is the X-ray image.

To obtain a good X-ray image, each component of the SYSTEM plays a

determining role and must be appropriately managed and integrated with all the

others.

No system is able to provide results fulfilling the requirements of any existing

Quality Control protocols without an appropriate calibration that takes into

consideration all the specific characteristics of each single components.

The aim of this procedure is to integrate the system components for obtaining

the requested Optical Density and linearity over the widest range of breast

density.

1.2	Selecting film/screen combination

Some key information will be provided below on how to obtain a good image with minimum dose.

For maximum image quality, it is advisable to use only film/screen combinations

recommended by the manufacturer.

The United States regulations have made this recommendation a must.

The X-ray cassettes for mammography use have only one intensifying screen

that converts the X-rays to light that impresses the image on film.

Every screen has its own x-ray/light conversion factor that must be suited to the

film to be used.

Films for mammography use are single emulsion and have various sensitivities

and contrasts.

An appropriate choice of these products is reflected in image quality.

We leave this responsibility to the user, dutifully pointing out the difference that

it can represent in terms of x-ray Dose delivered to the patient.

The tables shows mAs for the same Optical Density on 5.5 cm of Plexiglas,

using two different films/screens.

The difference is considerable: 480/190=2.5 times!

MEDICAL SCIENTIFIC

Chapter 3 Page 24 of 90

МАММО-ЗМТ TECHNICAL MANUAL

1.3 Film Processor

Film for mammography use has treatment characteristics that are very specific

and different from those used for other film.

A dedicated Processor is highly recommended for treating mammography film,

which guarantees temperature and development time suited to the specific film.

Development temperature is generally between 34 and 35 °C and has a

marked influence on the radiological dose necessary to obtain the required

O.D.

MEDICAL SCIENTIFIC

Chapter 3 Page 25 of 90

МАММО-ЗМТ TECHNICAL MANUAL

2 - PROCEDURE PREPARATION

2.1	Software installation

Have software installed and serial port connection checked before starting the procedure.

2.2	Developer

The developer must be in perfect conditions for use with developing and fixing replaced a day beforehand.

2.3	X-ray cassette for tests

Select and mark a cassette that will be the only one used for all the tests.

If possible, select the cassette with optical density as close as possible to the

average measured on all those available.

2.4	Film pack for tests

Make sure that there is enough film in the pack being used to complete the procedure, otherwise open a new pack.

2.5	Test phantom

A Plexiglas test phantom is supplied with each unit, made up of three 2 cm

layers, one 1 cm layer and one 0.5 cm layer.

In each layer there is a small steel sphere placed at an increasing distance

from the left edge.

These spheres identify the layers in order to guarantee that the position and the

composed layers can be repeated and guarantee the best repeatability of x-ray

attenuation.

2.6	O.D. reference X-ray

Expose a film in Manual Mode 28kV technique, with a 4 cm phantom with optical density between 1 and 2.

If during the procedure Processor stability is in doubt, repeat the same X-ray to quickly estimate O.D. repeatability.

2.7	Time for carrying out procedure

For best results, it is recommended to carry it out without interruptions for lunch or other.

2.8	Anode filter combination

If the unit has an automatic Mo/Rh filter manually select the correct Fix configuration from the control panel in order to avoid automatic selection of filter or anode at higher phantom thickness.

MEDICAL SCIENTIFIC

Chapter 3 Page 26 of 90

МАММО-ЗМТ TECHNICAL MANUAL

2.9 Detector calibration

Detector sensitivity is of maximum importance in order to guarantee the correct

dynamic and correspondence of the exposure parameters.

To facilitate control and calibration of parameter [r] each unit has a plain block

of Plexiglas (phantom) for a rapid control and a simple periodical re-calibration.

The phantom is a relative reference that is characterized vs. an absolute

reference during factory calibration.

The phantom is an integrating part of the unit, can be used also by the operator

and must always be available to the operator throughout time.

MEDICAL SCIENTIFIC	Chapter 3 Page 27 of 90

For [r] testing, the system must be configured as shown:

МАММО-ЗМТ TECHNICAL MANUAL

- Install Grid table 18x24

- Install 18x24 collimation plate

- Select LARGE focus

- Remove compression plate

- Remove cassette

- Select ZERO POINTS technique

- Disable potter bucky by inhibiting compressor release at exposure end and pressing the "Compression +" key

- Position detector at first stop on patient side

- Position the 7-cm-thick Phantom so as to completely cover the detector.

- Repeat the test for each one of the three fields.

This   configuration   was   selected   in   order  to  eliminate   all   the   variable components that could be placed between the beam source and the detector.

Carry out an exposure and then check that [r] value is within ±10 of the value

from the phantom label.

At the same time check that [o] value is not 0 (zero) or higher than 12.

If out of range adjust according to the following points 10 and 11.

Normal operating condition is restored when cassette is inserted in the table.

2.10	Offset [o] Calibration

The value of the detector offset appears in the lower left of the display when

Test mode is selected.

The offset value depends on the temperature.

The value must be checked after at least 5 minutes from turning the unit on and

it should be 10±5 .

The system loses its linearity if the value reaches zero or exceeds 20.

When adjusting GAIN [r] or Offset [o], large variations of one of them can affect

the other, check both and eventually correct before to leave calibration.

2.11	Detector [r] Gain Calibration

Value of [r] depends on:

- Precision of the kV value

- Precision of the mA anodic current

- Detector gain calibration.	/ Before proceeding with gain calibration, make sure that the values referred to in points a) and b) are correct.

The value of [r] appears in the lower left of the display when Automatic Mode is selected.

MEDICAL SCIENTIFIC

Chapter 3 Page 28 of 90

МАММО-ЗМТ TECHNICAL MANUAL

The value must be checked after at least 5 minutes from turning the unit on and it should be the same as the value shown on the test phantom +/-10. Оoffset; GGain ([r] on the display).

Turning offset clockwise increases

Turning gain clockwise  increases

2.12 Connecting PC for calibration

To connect the PC to the Mammo-ЗМТ, insert the serial cable into the service socket on the back side of the mammograph.

Service socket

MEDICAL SCIENTIFIC

Chapter 3 Page 29 of 90

МАММО-ЗМТ TECFTNICAL MANUAL

2.13 Selection of test layers

For the best overall results it's important that Zero Point calibration is carried on

in the most accurate conditions.

The minimum procedure is based on three thicknesses, 2, 4, 6 cm.

Before starting the procedure with Films, some "dry" test (cassette inserted but

no film) must be performed to find the most appropriate phantom thicknesses

+/-5mm, the related value of [p] must be checked to see that it is as much close

to the center of a kV value, and not near the margin between two kV values.

If this point does not fall in the middle of the kV interval, there can be a

difference of 0.5 kV between successive exposures; this MUST be avoided.

For checking, refer to the specific range and kV steps of the technique currently

under calibration (STD, HC, LD, EXT), focal spot and anode/filter combination.

Use the 5mm  Plexiglas layer or even remove/insert the compression plate in

order to accommodate p (plog)

Repeat the test for all thicknesses that are to be used.

2.14	Other Techniques and operating modes

Every operating mode has its own independent calibration

2.15	Automatic Molybdenum Rhodium filter

If the unit is equipped with an automatic Mo/Rh filter, filter selection is automatic

and depends on breast density, verified during pre-exposure.

The TRIP POINT T can be programmed, but it is not advisable to change it

unless there are well-evaluated reasons based on the statistical readings based

on the last 1300 exposures.

For the Rhodium filter,    the calibration curve of the relative Anode Filter

combination must be programmed.

During calibration of this curve, bear in mind that the part of the curve used will

be more to the right, beyond the relative p value at the threshold point T.

MEDICAL SCIENTIFIC

Chapter 3 Page 30 of 90

МАММО-ЗМТ TECHNICAL MANUAL

A part of the curve to the left of the T point could nevertheless be used if the operator has manually selected the Rhodium Filter.

2.16	Modification of the reference Density after calibration

Each calibration curve and related reference O.D. can be corrected to get it lighter or darker by means of option (O) to add or subtract a percentage value to move the calibration curve up or down, giving rise to an increase or a decrease of O.D..

2.17	Maximum breast density limit

With the DEATA PLUS software, it is possible to define a maximum work limit (LOW p) defined as maximum breast density, to interrupt exposure after the sampling impulse, releasing a dose of less than 1 mAs. The limit is defined according to the specific characteristics of each single installation, in order to avoid releasing excessive doses, which could make the Film unusable as a consequence of unforeseeable operating errors, such as an undeclared silicone prosthesis. In this case the AEC ERROR alarm is generated.

2.18	Calibration concept

Deata Plus procedure is based on two films for each phantom thickness having O.D. Higher and Lower than selected reference value as described in the following drawing.

A

MEDICAL SCIENTIFIC

Chapter 3 Page 31 of 90

МЛММО-ЗМТ TECHNICAL MANUAL

Phantom thickness for test are minimum 3.

Automatic data collection, interpolation and curve smoothing are provided,

operator has only to measure and input O.D.

Automatic procedure is also provided to generate Manual Density Steps after

that Zero Point curve has been generated by means of automatic procedure.

One point mode calibration is the last operation and can be performed for kV

and Phantom Thickness specified in the applicable protocols.

2.19	Linearity check

After calibration, linearity check must be performed to verify that all what you

did was correct.

After setting the Manual Density command to ZERO, make one exposure for

each phantom thickness and kV specified in the quality protocol

If something needs to be fine adjusted, tools are provided for manual calibration

of specific points.

2.20	Unused Operating Technique and Film Screen

In order to avoid operating errors and misunderstandings to the operator, if only one Film Screen and Operating Technique, STD, EXT, HC, LD has been calibrated, all the others can be disabled by means of DIP 5 SW1 PCB 01-170. If more than one Film Screen and Operating Technique are calibrated and available for normal use, the above function can't be used, refer to next paragraph.

2.21	Disabling specific Operating Technique and/or Film Screen

Disabling an Operating Technique or the entire Film Screen is done by storing the relative curve or curves with the value of i=0 for the first pair of values that correspond to p = 0.

Tempted exposure with Techniques with a curve that has i=0 for p=0 does not start tube rotation and generates the message "IN-ACTIVE"

MEDICAL SCIENTIFIC

Chapter 3 Page 32 of 90

МАММО-ЗМТ TECHNICAL MANUAL

3 - CALIBRATION PROCEDURE

Two different online Help are available ♦ Help for software functions

♦ Open or Create calibration file as described in Deata Plus Help

♦ Verify that appropriate Deata Type is defined according to the configuration of the specific unit you have to calibrate.

♦ Take note of Environmental Conditions for future reference

MEDICAL SCIENTIFIC

Chapters Page 33 of 90

МАММО-ЗМТ TECHNICAL MANUAL

♦ Process and record data from a sensitometric strip

Zero point calibration curve      I

♦ Configure Mammo unit for the Operating Technique to calibrate.

Take an exposure to check that a calibration curve is already stored in the unit,

so that the exposure parameters appear in the "Emission data" window.

If "Technique not programmed" alarm comes out, default calibration curve must

be stored before to go ahead with calibration.

Repeat test exposure to check that calibration curve is now available.

Check for appropriate phantom thickness as described in chapter "Procedure

Preparation", paragraph "Selection of test layers".

♦ Select Reference O.D. for calibration

Reference O.D. can also be selected later from inside Zero point curve procedure.

calibration

MEDICAL SCIENTIFIC

Chapter 3 Page 34 of 90

МАММО-ЗМТ TECHNICAL MANUAL

3.1 Zero Point calibration

♦ Select Zero Point calibration curve

and start procedure from one of the phantom thicknesses having  Manual Density set to ZERO on the control panel of Mammo unit.

Two different calibration possibilities are available: Automatic and Manual

Automatic mode is recommended as most powerful tool giving access to other

automatic functions of this calibration software.

Manual calibration is only provided for expert technicians to perform special

tasks.

Automatic

exposure parameters

are automatically recorded in the table

MEDICAL SCIENTIFIC

Chapter 3 Page 35 of 90

МЛММО-ЗМТ TECHNICAL MANUAL

after keyboard input of measured O.D.

Parameters are recorded as High values if O.D. is higher than reference and

vice versa.

If first O.D. for a given thickness il Low, next exposure must have Manual

density set to +1 or more depending on how much difference exists between

Reference O.D. and last exposure.

If O.D. was higher than reference, set Manual Density to -1 or more.

When High and Low values are complete change to the next thickness and

repeat the procedure.

A minimum of two phantom thickness is necessary before to get the Change

button green, but for the best results a minimum of three phantom thickness is

required.

After clicking the Change button, calibration curve is automatically reshaped.

To accept click Save to File and transfer new curve to the unit.

Values in the table are saved to file and can be recalled by means of Reload Last button

MEDICAL SCIENTIFIC

Chapter 3 Page 36 of 90

МАММО-ЗМТ TECHNICAL MANUAL

to complete or modify calibration curve.

Important NOTE

If default curve used for calibration is too much different from characteristics of Film/screen to be calibrated, may happen that manual density adjustment is not wide enough to obtain High and Low densities for each phantom thickness. In such circumstance it's necessary to modify the default curve and store it again in the unit before to proceed with next points/thickness. Calibration curve can be modified editing points manually or by means of %O.D. correction

the same concept remains valid for next Manual procedure.

MEDICAL SCIENTIFIC

Chapter 3 Page 37 of 90

МАММО-ЗМТ TECHNICAL MANUAL

Manual

Exposure parameters

have to be manually recorded into calibration forms,

processed by means of Log Interpolator,

MEDICAL SCIENTIFIC

Chapter 3 Page 38 of 90

МАММО-ЗМТ TECHNICAL MANUAL

to obtain coordinates Plog (p) and Pulse (i) to build up the calibration curve.

In Manual calibration procedure curve shape must be manually adjusted for plog = 0 and plog = 159 being the extremities of the curve.

♦ For each phantom thickness used for calibration, make a test exposure and check that O.D. is now within requested tolerance.

MEDICAL SCIENTIFIC

Chapter 3 Page 39 of 90

МЛММО-ЗМ1 TECHNICAL MANUAL

3.2 Manual Density steps

select automatic

♦ Select Manual density control

define O.D. step change

and click OK. Table will be automatically changed.

Save to file and transfer new data to the unit

Check with exposures that step O.D. changes according to selected value.

Chapter 3 Page 40 of 90

I

МАММО-ЗМТ TECHNICAL MANUAL

3.3 One point calibration

i.6 une point calibration

n One point mode kV are selected by the operator and for a given thickness

hfiv пяп hp>

they can be

<	than kV in zero point mode (dkV between -14 and -1), or
> than kV in zero point mode (dkV between +1 and +10).

If they are > than kV in zero point mode, they can also be:

<	than 30 kV
>= of 30kV.

For each specific condition correction factors % are defined to compensate O.D. in function of kV difference between Zero (CPU) and One point (operator).

Select %

Blect % value to change in the appropriate column. If after exposure, parameter from communication window was +2 , One point was 2 kV high than Zero point kV

d

er

If measured O.D. was outside tolerance, % can be adjusted with the same concept of High and Low O.D. as seen during the Zero point calibration.

For easy operation, Log Interpolator is available

ffi. One port interpolator

MEDICAL SCIENTIFIC

Chapter 3 Page 41 of 90

МАММО-ЗМТ TECHNICAL MANUAL

values must be filled manually, when clicking OK, % in the selected cell will be automatically adjusted.

Repeat the same procedure for all points with: fix phantom and variable kV, fix kV and variable thickness according to the next forms.

MRDICAI. SCIENTIFIC	Chapter 3 Page 42 of 90

МАММО-ЗМТ TECHNICAL MANUAL

3.4 Rhodium filter

If automatic Mo/Rh filter is installed, Rhodium  Zero Point curve and One point

correction factors have to be defined as before for Mo/Mo.

As Mo/Rh is used for large breasts, only the upper part of calibration curve is

relevant.

The same will be for correction factors of One point mode.

For calibration, from the control panel, select Mo/Rh FIX.

For normal operation, Rhodium filter is automatically selected depending on

breast density.

If necessary (not recommended), Trip Point can be adjusted by means of :

If number is decreased Rh filter will come in for larger breasts and vice versa. The concept of this paragraph is extended if a double trace tube is present.

3.5 Maximum breast density limit

Select Zero Point calibration curve

One Point calibration curve

Click inside Trip Point Low [p]

И Trip Point Low И   [0

and define a new value.

If [p] value will be lower than the defined limit, " Breast too dense" alarm will come on and exposure is stopped releasing only the dose of pre-exposure x-ray pulse.

3.6 AEC calibration files backup

For future reference is essential that CD with Calibration software and AEC Calibration files remains available at the site.

AEC calibration files backup can be done on a floppy disk or added to the same software CD if PC used for calibration has CD writer and burning software.

MEDICAL SCIENTIFIC	Chapter 3 Page 43 of 90