## Alarm 122
Code 122 refers to the transmission of the longitudinal drawer and in particular to activation 24S6.
The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

Events that may cause the error	Problem solution
Activation not powered or powered by an out of range voltage value / LED HV: OFF / LED FAU: ON	Check tracks on board 24S9, the operation of power supply board 24S5, fuses on board 24S5, the power supplies from the power transformer, connections to transformer, fuses on transformer, the power supplies of the transformer and line voltage. / Replace board 24S9, 24S6, fuses or board 24S5, restore connections, replace fuses or power transformer 22TR1.

Events that may cause the error	Problem solution
Activation powered by an out of range voltage value / LED HV: OFF / LED FAU: ON	Check line voltage, the operation of power supply board 24S5, the power supplies from the power transformer, the power supplies of the transformer. Replace board 24S5 and power transformer 22TR1.
Activation blocked by thermal protection / LED TER: ON / LED FAU: ON	Ventilators off or not functional, use of continuous transmission function. / To restore operation switch activation off, let it cool down and switch the equipment back on.
Motor short-circuit / LED HV: ON / LED FAU: ON	Check tracks on board 24S9, the activation/motor tie cable and motor 3MT11. / Replace cable and motor.
Error in motor connection / LED HV: ON / LED FAU: ON	Refer to functional diagrams to restore connections.
Tracks interruptions on board 24S9 / Input 25S2-IN3 off	Restore connections or replace board.
Interruption in flat-cable 24S9CFC2-25S2CN2 / Input 25S2-IN3 off	Replace flat-cable.
Rosin joints on the connectors of board 24S9 and 25S2 / Input 25S2-IN3 off	Restore joints or replace boards.