OPERA T

Universal remote-controlled table

TROUBLESHOOTING

Ver.1.1 - January 2004

Rev. n.01 27/01/2004

This manual is suitable for all versions of the universal remote-controlled table.

This page intentionally left blank

1	Troubleshooting	5

1.1	Aim	5

1.2	Alarms/messages display	5

1.3	Alarms/messages list	6

1.4	Alarm 001	8

1.5	Alarm 002	9

1.6	Alarm 003	12

1.7	Alarm 004	12

1.8	Alarm 011	13

1.9	Alarm 012	14

1.10	Alarm 013	17

1.11	Alarm 014	17

1.12	Alarm 015	18

1.13	Alarm 021	18

1.14	Alarm 022	20

1.15	Alarm 023	22

1.16	Alarm 024	23

1.17	Alarm 025	23

1.18	Alarm 026	24

1.19	Alarm 027	24

1.20	Alarm 028	25

1.21	Alarm 031	26

1.22	Alarm 032	27

1.23	Alarm 033	30

1.24	Alarm 034	30

1.25	Alarm 035	31

1.26	Alarm 036	31

1.27	Alarm 037	32

1.28	Alarm 041	32

1.29	Alarm 042	33

1.30	Alarm 051	35

1.31	Alarm 052	36

1.32	Alarm 061	38

1.33	Alarm 062	39

1.34	Alarm 071	41

1.35	Alarm 072	42

1.36	Alarm 081	44

1.37	Alarm 082	45

1.38	Alarm 091	47

1.39	Alarm 092	48

1.40	Alarm 096	49

1.41	Alarm 097	50

1.42	Alarm 098	51

1.43	Alarm 099	51

1.44	Alarm 101	53

1.45	Alarm 111	55

1.46	Alarm 121	56

1.47	Alarm 122	57

1.48	Alarm 131	59

1.49	Alarm 132	60

1.50	Alarm 133	62

1.51	Alarm 141	63

1.52	Alarm 142	64

1.53	Alarm 143	66

1.54	Alarm 151	67

1.55	Alarm 152	68

1.56	Alarm 161	70

1.57	Alarm 163	71

1.58	Alarm 164	72

1.59	Alarm 189÷196	72

1.60	Alarm 197	72

1.61	Alarm 198	73

1.62	Alarm 199	73

1.63	Alarm 200	73

1.64	Alarm 201	74

1.65	Alarm 202	74

Notes:

# Troubleshooting

## Aim

The aim of this document is to assist personnel in charge of testing or repairing the equipment in troubleshooting during machine set-up and repair.

This guide might not include all possible fault-related occurrences; however, it is intended to provide some guidelines on possible problems and their solutions. Moreover, it is understood that most boards/control activations, due to their complexity and construction type, can be repaired with suitable equipments/materials at the board testing unit.

For further information on troubleshooting it is recommended the use of <Telediagnosis>. Refer to manual “Opera T – Telediagnosis”.

## Alarms/messages Display

Alarms and messages are shown as numerical codes on the console display.

## Alarms/messages list

| TRANSMISSION | ERROR CODE | DESCRIPTION |
| --- | --- | --- |
| TILTING | 001 | Potentiometer |
|  | 002 | Motor’s board FAULT signal |
|  | 003 | Inverted motor or dynamo or DAC |
|  | 004 | Gear box reversibility |
| ELEVATION | 011 | Potentiometer |
|  | 012 | Motor’s board FAULT signal |
|  | 013 | Inverted motor or dynamo or DAC |
|  | 014 | Gear box reversibility |
|  | 015 | I.I. / tilting curve mistake (software limits only for T90e) |
| TUBE SUPPORT (COLUMN) | 021 | Potentiometer |
|  | 022 | Motor’s board FAULT signal |
|  | 023 | Inverted motor or dynamo or DAC during scanning |
|  | 024 | Inverted motor or dynamo or DAC during incidence or centering |
|  | 025 | Inverted motor or dynamo or DAC in tomography |
|  | 026 | Inverted motor or dynamo or DAC in scanning + incidence/centering |
|  | 027 | Gear box reversibility |
|  | 028 | Column over the maximum incidence limit (40°) |
| SPOT FILM DEVICE | 031 | Potentiometer |
|  | 032 | Motor’s board FAULT signal |
|  | 033 | Inverted motor or dynamo or DAC during scanning |
|  | 034 | Inverted motor or dynamo or DAC in incidence/centering |
|  | 035 | Inverted motor or dynamo or DAC in tomography |
|  | 036 | Inverted motor or dynamo or DAC in scanning+ incidence/centering |
|  | 037 | Gear box reversibility |
| FOCAL DISTANCE | 041 | Zero input alarm |
|  | 042 | Motor’s board FAULT signal |
| PARALLEL COLLIMATOR | 051 | Potentiometer |
|  | 052 | Motor problem (blades didn’t reach position, check clutch) |
| PERPENDICULAR COLLIMATOR | 061 | Potentiometer |
|  | 062 | Motor problem (blades didn’t reach position, check clutch) |

| TRANSMISSION | ERROR CODE | DESCRIPTION |
| --- | --- | --- |
| IRIS COLLIMATOR | 071 | Potentiometer |
|  | 072 | Motor problem (blades didn’t reach position, check clutch) |
| PARALLEL DIAPHRAGM | 081 | Zero input alarm |
|  | 082 | Motor’s board FAULT signal |
| CROSS DIAPHRAGM | 091 | Zero input alarm |
|  | 092 | Motor’s board FAULT signal |
| PARALLEL JAWS | 101 | Potentiometer |
| PERPENDICULAR JAWS | 111 | Potentiometer |
| IN/OUT CASSETTE TRAY | 121 | Zero input alarm |
|  | 122 | Motor’s board FAULT signal |
| TRANSVERSE CASSETTE TRAY (cross division only) | 131 | Zero input alarm |
|  | 132 | Motor’s board FAULT signal |
|  | 133 | Central drawer input alarm |
| I.I. | 141 | Zero input alarm (I.I. up) |
|  | 142 | Motor’s board FAULT signal |
|  | 143 | Low IB input alarm |
| GRID | 151 | Zero input alarm |
|  | 152 | Motor’s board FAULT signal |
| LONGITUDINAL TABLE TOP | 161 | Potentiometer |
|  | 022 | Motor’s board FAULT signal |
|  | 163 | Inverted motor or dynamo or DAC |
|  | 164 | Gear box reversibility |
| VARIOUS | 097 | A command is active when Power ON |
|  | 098 | Incorrect transmission/optical fiber broken |
|  | 099 | Power OFF |
|  | 197 | ASSI board problem |
|  | 198 | Transmission error on RS232 in installation |
|  | 199 | Character transmission error on RS232 |
|  | 200 | EEPROM absent/broken |
|  | 201 | EEPROM virgin |
|  | 202 | Alarm not managed |

## Alarm 001

Code 001 refers to the transmission of the tilt and in particular to the position transducer (potentiometer and related circuits).

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Possible alarm causes and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Potentiometer 3PT1 fault H | Replace potentiometer. Reconnect, reacquire and check limits following setup procedure. |
| Potentiometer 3PT1 exceeds acquisition limits L | Bring tilt within defined limits, following setup procedure or manually. Recheck control limits. |
| Tie cable disconnected between board 25S3 and 3PT1 VL | Reconnect cable. CAUTION : ensure shield of cable is connected to grounding ring. |
| Interruption in connection cable between board 25S3 and 3PT1 L | Replace cable. |
| Incorrect connections on cable between board 25S3 and 3PT1 (only during quality controls) | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 25S3 VVL | Restore connections or replace board. |
| Cold welding on the connectors of board 25S3 (only during quality controls) | Restore joints or replace board. |
| Interruption in flat-cable 25S3CN2-25S1CN2 (only during quality controls) | Replace flat-cable. |
| Fault on reading circuit ADC1 of board 25S1 CPU VVL | Replace board. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| +/-15V failure on power supplies of reading circuit ADC1 in board 25S1 VVL | Check fuses on board 25S1, the power supplies and connections from the service transformer, fuses and power supplies on transformer. Replace fuses or board 25S1, restore connections, replace fuses or auxiliaries transformer 22TR2. |
| Sporadic errors L | Error caused by external interferences. To restore activation switch the equipment off and on. |

## Alarm 002

Code 002 refers to the transmission of the tilt and in particular to activation 23S1.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Board 23S1 absent or not functional LED OK: OFF H | While equipment is switched off insert or replace board. |
| Activation powered by an out of range voltage value LED OK: OFF VL | Check board 23S9 and especially fuses, the power supplies and connections from the power transformer, the power supplies of the transformer and line voltage. Replace fuses or board 23S9, restore connections, replace fuses or power transformer 22TR1. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Inverted motor connection LED O.K.: OFF LED m.d.: ON L (during quality controls) | Invert motor connection. |
| Motor disconnected LED O.K.: OFF LED O.C.: ON (possible) | Refer to functional diagrams to restore connections. |
| Motor short-circuit LED O.K.: OFF LED O.C.: ON (possible) | Check Motor, the generator/activation tie cable, board 23S9 and activation 23S1. Replace Motor Reducer, tie cable, boards 23S9 and 23S1. |
| Inverted dynamo connection LED O.K.: OFF LED m.d.: ON VVL (during quality controls) | Invert generator connection. |
| Dynamo disconnected LED O.K.: OFF LED m.d.: ON L | Refer to functional diagrams to restore connections. |
| Dynamo short-circuit LED O.K.: OFF LED m.d.: ON VVL | Check generator, the generator/activation tie cable, board 23S9 and activation 23S1. Replace Motor Reducer, tie cable, boards 23S9 and 23S1. |
| Board fault by thermal protection LED O.K.: OFF LED S.T.: ON VL | Continuous use of transmission function. To restore operation switch activation off, let it cool down and switch it back on. |
| Tracks interruptions on board 23S9 Input 25S3-IN21 off Output 25S3-OUT14 off Dac VAN3 out of range VVL | Restore connections or replace board. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Interruption in flat-cable 23S9CFC1-25S3CFC1 Input 25S3-IN21 off Output 25S3-OUT14 off Dac VAN3 out of range (only during quality controls) | Replace flat-cable. |
| Cold welding on the connectors of board 23S9 and 25S3 Input 25S3-IN21 off Output 25S3-OUT14 off Dac VAN3 out of range VVL (only during quality controls) | Restore joints or replace boards. |
| Input 25S3-IN21 not functional Output 25S3-OUT14 not functional Dac VAN3 out of range VVL | Refer to testing manual to check with Test SW. Replace board I/O 25S3, CPU 25S1 and flat-cables. |
| Power supply failure +24Vcc on board 25S3 for digital input/output VL | Check connections, power supplies on boards 25S3, the power supplies from the service transformer, fuses and power supplies on transformer. Replace board 25S3, connections, fuses or auxiliaries transformer. |
| Power supply failure +5Vcc or +/-12Vcc on board 25S1 for digital input/output VL | Check connections, power supplies and fuses on boards 25S1, the power supplies from the service transformer, fuses and power supplies on transformer. Replace board 25S1, connections, fuses or auxiliaries transformer. |

## Alarm 003

Code 003 refers to the transmission of the tilt during movement. Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of the movement.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Inverted motor and dynamo VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| Inverted DAC connections VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| All events that may cause error 001 (problems with potentiometer and related circuits) | See error 001. (can be the potentiometer) |

## Alarm 004

Code 004 refers to the transmission of the tilt and in particular to the motor reducer.

Through the position transducer (potentiometer and related circuits) the software determines the movement without command of the transmission.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Motor gear box reversibility VL | Replace motor-gear box |
| All events that may cause error 001 (problems with potentiometer and related circuits) | See error 001. |

## Alarm 011

Code 011 refers to the transmission of the elevation and in particular to the position transducer (potentiometer and related circuits).

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Potentiometer 3PT2 fault H | Replace potentiometer. Reconnect, reacquire and check limits following set up procedure. |
| Potentiometer 3PT2 exceeds acquisition limits L | Bring elevation within limits, following setup procedure or manually. Recheck control limits. |
| Tie cable disconnected between board 25S3 and 3PT2 VL | Reconnect cable. CAUTION: ensure shield of cable is connected to grounding ring |
| Interruption in connection cable between board 25S3 and 3PT2 L | Replace cable. |
| Incorrect connections in tie cable between board 25S3 and 3PT2 (only during quality controls) | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 25S3 VVL | Restore connections or replace board. |
| Cold welding on the connectors of board 25S3 (only during quality controls) | Restore joints or replace board. |
| Interruption in flat-cable 25S3CN2-25S1CN2 (only during quality controls) | Replace flat-cable. |
| Fault on reading circuit ADC2 of board 25S1 CPU VVL | Replace board. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Power supply failure +/-15V of reading circuit ADC2 of board 25S1 VVL | Check fuses on board 25S1, the power supplies from the service transformer, the connections, fuses and power supplies of the transformer. Replace fuses or board 25S1, restore connections, replace fuses or auxiliaries transformer. |
| Sporadic errors L | Error caused by external interferences. To restore activation switch the equipment off and on. |

## Alarm 012

Code 012 refers to the elevation transmission and in particular to activation 23S2.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Activation 23S2 absent or not functional LED OK: OFF H | While equipment is switched off insert or replace board. |
| Activation powered by an out of range voltage value LED OK: OFF VL | Check board 23S9 and especially fuses, the power supplies and connections from the power transformer, power supplies and line voltage. Replace fuses or board 23S9, restore connections, replace fuses or power transformer 22TR1. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Inverted motor connection LED O.K.: OFF LED m.d.: ON L (during quality controls) | Invert motor connection. |
| Motor disconnected LED O.K.: OFF LED O.C.: ON (possible) | Refer to functional diagrams to restore connections. |
| Motor short-circuit LED O.K.: OFF LED O.C.: ON (possible) | Check Motor, the generator/activation tie cable, board 23S9 and activation 23S2. Replace Motor Reducer, tie cable, board 23S9 and 23S2. |
| Inverted dynamo connector LED O.K.: OFF LED m.d.: ON VVL (during quality controls) | Invert generator connection. |
| Dynamo disconnected LED O.K.: OFF LED m.d.: ON L | Refer to functional diagrams to restore connections. |
| Dynamo short-circuit LED O.K.: OFF LED m.d.: ON VVL | Check generator, the generator/activation tie cable, board 23S9 and activation 23S2. Replace Motor Reducer, tie cable, board 23S9 and 23S2. |
| Board fault by thermal protection LED O.K.: OFF LED S.T.: ON VL | Use of continuous transmission function. To restore operation switch activation off, let it cool down and switch it back on. |
| Tracks interruptions on board 23S9 Input 25S3-IN22 off Output 25S3-OUT20 off Dac VAN4 out of range VVL | Restore connections or replace board. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Interruption in flat-cable 23S9CFC1-25S3CFC1 Input 25S3-IN22 off Output 25S3-OUT20 off Dac VAN4 out of range (only during quality controls) | Replace flat-cable. |
| Cold welding on the connectors of board 23S9 and 25S3 Input 25S3-IN22 off Output 25S3-OUT20 off DAC VAN4 out of range VVL (only during quality controls) | Restore joints or replace boards. |
| Input 25S3-IN22 not functional Output 25S3-OUT20 not functional Dac VAN4 out of range VVL | Refer to testing manual to check with Test SW. Replace board I/O 25S3, CPU 25S1 and flat-cables. |
| Power supply failure +24Vcc on board 25S3 for digital input/output VL | Check connections, power supplies on boards 25S3, the power supplies from the service transformer, fuses on transformer and power supplies. Replace board 25S3, restore connections, replace fuses or auxiliaries transformer. |
| Power supply failure +5Vcc or +/-12Vcc on board 25S1 for digital input/output VL | Check connections, power supplies on boards 25S1, the power supplies from the service transformer, fuses and power supplies on transformer. Replace boards 25S1, restore connections, replace fuses or auxiliaries transformer |

## Alarm 013

Code 013 refers to the transmission of the elevation during movement. Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of the movement.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Inverted motor and dynamo VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| Inverted DAC connections VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| All events that may cause error 011 (problems with potentiometer and related circuits) | See error 011. (can be the potentiometer) |

## Alarm 014

Code 014 refers to the transmission of the elevation and in particular to the motor reducer. Through the position transducer (potentiometer and related circuits) the software determines the movement without command of the transmission.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Motor gear box reversibility | Replace motor -gear box |
| All events that may cause error 011 (problems with potentiometer and related circuits) | See error 011. |

## Alarm 015

Code 015 refers to the transmission of the elevation during the synchronised movement with the tilt.

Through the position transducers (potentiometers and related circuits) the software determines the incorrect direction of the movements.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Inverted motor and dynamo due to tilt and elevation VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| --- | --- |
| Inverted DAC connections due to tilt and elevation VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| All events that may cause error 001 and 011 (problems with potentiometer and related circuits) | See error 001 and 011. |

## Alarm 021

Code 021 refers to the transmission of the tube support and in particular to the position transducer (potentiometer and related circuits).

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Potentiometer 3PT3 fault H | Replace potentiometer. Reconnect, reacquire and check control stops following setup procedure. |
| Potentiometer 3PT3 exceeds acquisition limits L | Bring transmission within defined limits, following setup procedure or manually and recheck control limits. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Interruption in connection cable between board 25S3 and 3PT3 L | Replace cable. |
| Incorrect connections in tie cable between board 25S3 and 3PT3 (only during quality controls) | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 25S3 VVL | Restore connections or replace board. |
| Cold welding on the connectors of board 25S3 (only during quality controls) | Restore joints or replace board. |
| Interruption in flat-cable 25S3CN2-25S1CN2 (only during quality controls) | Replace flat-cable. |
| Fault on reading circuit ADC3 of board 25S1 CPU VVL | Replace board. |
| Power supply failure +/-15V of reading circuit ADC3 of board 25S1 VVL | Check fuses on board 25S1, the power supplies from the service transformer, the connections, fuses and power supplies of the transformer. Replace fuses or board 25S1, restore connections, replace fuses or auxiliaries transformer 22TR1. |
| Sporadic errors L | Error caused by external interferences. To restore activation switch the equipment off and on. |

## Alarm 022

Code 022 refers to the transmission of the tube support and in particular to activation 23S3.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Activation 23S3 absent or not functional LED OK: OFF H | While equipment is switched off insert or replace aboard. |
| Activation powered by an out of range voltage value LED OK: OFF VL | Check board 23S9 and especially fuses, the power supplies and connections from the power transformer, power supplies and line voltage. Replace fuses or board 23S9, restore connections, replace fuses or power transformer 22TR1. |
| Inverted motor connection LED O.K.: OFF LED m.d.: ON L (during quality controls) | Invert motor connection. |
| Motor disconnected LED O.K.: OFF LED O.C.: ON (possible) | Refer to functional diagrams to restore connections. |
| Motor short-circuit LED O.K.: OFF LED O.C.: ON (possible) | Check Motor, the generator/activation tie cable, board 23S9 and activation 23S3. Replace Motor Reducer, tie cable, boards 23S9 and 23S3. |
| Inverted dynamo connection LED O.K.: OFF LED m.d.: ON VVL (during quality controls) | Invert dynamo connection. |
| Disconnected dynamo LED O.K.: OFF LED m.d.: ON L | Refer to functional diagrams to restore connections. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Dinamo short-circuit LED O.K.: OFF LED m.d.: ON VVL | Check generator, the generator/activation tie cable, board 23S9 and activation 23S3. Replace Motor Reducer, tie cable, boards 23S9 and 23S3. |
| Board fault by thermal protection LED O.K.: OFF LED S.T.: ON VL | Use of continuous transmission function. To restore operation switch equipment off, let it cool down and switch it back on. |
| Tracks interruptions on board 23S9 Input 25S3-IN11 off Output 25S3-OUT13 off DAC VAN1 out of range VVL | Restore connections or replace board. |
| Interruption in flat-cable 23S9CFC1-25S3CFC1 Input 25S3-IN11 off Output 25S3-OUT13 off DAC VAN1 out of range (only during quality controls) | Replace flat-cable. |
| Cold welding on the connectors of board 23S9 and 25S3 Input 25S3-IN11 off Output 25S3-OUT13 off DAC VAN1 out of range VVL (only during quality controls) | Restore joints or replace boards. |
| Input 25S3-IN11 not functional Output 25S3-OUT13 not functional DAC VAN1 out of range VVL | Refer to testing manual to check with Test SW. Replace board I/O 25S3, CPU 25S1 and flat-cable. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Power supply failure +24Vcc for input/output board 25S3 VL | Check connections, power supplies on boards 25S3, the power supplies from the service transformer, fuses on transformer and power supplies. Replace board 25S3, restore connections, replace fuses or auxiliaries transformer. |
| Power supply failure +5Vcc or +/-12Vcc for input/output board 25S1 VL | Check connections, power supplies and fuses on boards 25S1, the power supplies from the service transformer, fuses and power supplies on transformer. Replace board 25S1, restore connections, replace fuses or auxiliaries transformer |

## Alarm 023

Code 023 refers to the transmission of the tube support during movement in dynamic area. Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Inverted motor and dynamo VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| Inverted DAC connections VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| All events that may cause error 021 (problems with potentiometer and related circuits) | See error 021. (can be the potentiometer) |

## Alarm 024

Code 024 refers to the transmission of the tube support during movement in angling/centering. Through the position transducer (potentiometer and related circuits) the  software defines the incorrect direction of movement.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability | Problem solution |
| --- | --- |
| Inverted motor and dynamo VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| Inverted DAC connections VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| All events that may cause error 021 (problems with potentiometer and related circuits) | See error 021. |

## Alarm 025

Code 025 refers to the transmission of the tube support during movement in tomography. Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Inverted motor and dynamo VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| Inverted DAC connections VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| All events that may cause error 021 (problems with potentiometer and related circuits) | See error 021. |

## Alarm 026

Code 026 refers to the transmission of the tube support during simultaneous movement of dynamic area and angling/centering. Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Inverted motor and dynamo VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| Inverted DAC connections VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| All events that may cause error 021 (problems with potentiometer and related circuits) | See error 021. |

## Alarm 027

Code 027 refers to the transmission of the tube support and in particular to the motor reducer. Through the position transducer (potentiometer and related circuits) the software determines the movement without command of the transmission.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Motor-gear box reversibility VL | Replace motor-gear box. |
| All events that may cause error 021 (problems with potentiometer and related circuits) | See error 021. |

## Alarm 028

Code 028 refers to the transmission of the tube support and in particular it is activated when there is an angle with the serial radiographic design higher than 42°.

Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement, to avoid the disengagement of the rod.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the layer display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Inverted motor and dynamo tube support or SFD VVL (only during quality controls) | Refer to functional diagrams to restore connections. |
| Inverted DAC connections tube support or SFD | Refer to functional diagrams to restore connections. |
| All events that may cause errors 021-031 (problems with potentiometers and related circuits) | See errors 021-031. |

## Alarm 031

Code 031 refers to the transmission of the serial radiographic design and in particular to the position transducer (potentiometer and related circuits).

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Potentiometer 3PT4 fault H | Replace potentiometer. Reconnect, reacquire and check limits following set up procedure |
| Potentiometer 3PT4 exceeds acquisition limits L | Bring transmission within defined limits, following setup procedure or manually and recheck control limits. |
| Tie cable disconnected between board 25S3 and 3PT4 VL | Reconnect cable. CAUTION : ensure shield of cable is connected to grounding ring. |
| Interruption in connection cable between board 25S3 and 3PT4 L | Replace cable. |
| Incorrect connections in tie cable between board 25S3 and 3PT4 (only during quality controls) | Refer to functional diagrams to restore connections. . |
| Tracks interruptions on board 25S3 VVL | Restore connections or replace board. |
| Cold welding on the connectors of board 25S3 (only during quality controls) | Restore joints or replace board. |
| Interruption in flat-cable 25S3CN2-25S1CN2 (only during quality controls) | Replace flat-cable. |
| Fault on reading circuit ADC4 of board 25S1 CPU VVL | Replace board. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Power supply failure +/-15V reading circuit ADC4 of board 25S1 | Check fuses on board 25S1, the power supplies from the service transformer, the connections, fuses and power supplies of the transformer. Replace fuses or board 25S1, restore connections, replace fuses or auxiliaries transformer. |
| Sporadic errors | Error caused by external interferences. To restore activation switch the equipment off and on. |

## Alarm 032

Code 032 refers to the transmission of the serial radiographic design and in particular to activation 23S4.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Board 23S4 absent or not functional LED OK: OFF H | While equipment is switched off insert or replace board. |
| Activation powered by an out of range voltage value LED OK: OFF VL | Check board 23S9, fuses on board 23S9, the power supplies from the power transformer, the connections and power supplies of the transformer, line voltage. Replace fuses or board 23S9, restore connections, replace fuses or power transformer 22TR1. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Inverted motor connection LED O.K.: OFF LED m.d.: ON L (during quality controls) | Invert motor connection. |
| Motor disconnected LED O.K.: OFF LED O.C.: ON (possible) | Refer to functional diagrams to restore connections. |
| Motor short-circuit LED O.K.: OFF LED O.C.: ON (possible) | Check Motor, the generator/activation tie cable, board 23S9 and activation 23S4. Replace Motor Reducer, tie cable, board 23S9 and 23S4. |
| Inverted dynamo connection LED O.K.: OFF LED m.d.: ON VVL (only during quality controls) | Invert dynamo connection. |
| Dynamo disconnected LED O.K.: OFF LED m.d.: ON VVL | Refer to functional diagrams to restore connections. |
| Generator short-circuit LED O.K.: OFF LED m.d.: ON | Check generator, the generator/activation tie cable, board 23S9 and activation 23S4. Replace Motor Reducer, tie cable, board 23S9 and 23S4. |
| Activation blocked by thermal protection LED O.K.: OFF LED S.T.: ON | Use of continuous transmission function. To restore operation switch equipment off, let activation cool down and switch equipment back on. |
| Tracks interruptions on board 23S9 Input 25S3-IN12 off Output 25S3-OUT12 off Dac VAN2 out of range | Restore connections or replace board. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Interruption in flat-cable 23S9CFC1-25S3CFC1 Input 25S3-IN12 off Output 25S3-OUT12 off Dac VAN2 out of range | Replace flat-cable. |
| Rosin joints on the connectors of board 23S9 and 25S3 Input 25S3-IN12 off Output 25S3-OUT12 off Dac VAN2 out of range | Restore joints or replace boards. |
| Input 25S3-IN12 not functional Output 25S3-OUT12 not functional Dac VAN2 out of range | Refer to testing manual to check with Test SW. Replace board I/O 25S3, CPU 25S1 and flat-cables. |
| Power supply failure +24Vcc for input/output board 25S3 | Check connections, power supplies on boards 25S3, the power supplies from the service transformer, fuses and power supplies on transformer. Replace boards 25S3, restore connections, replace fuses or auxiliaries transformer 22TR1. |
| Power supply failure +5Vcc or +/-12Vcc for input/output board 25S1 | Check connections, power supplies on boards 25S1, fuses on board 25S1, the power supplies from the service transformer, fuses on transformer. Replace board 25S1, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 033

Code 033 refers to the transmission of the serial radiographic design during movement in dynamic area. Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Inverted motor and generator | Refer to functional diagrams to restore connections. |
| Inverted dac connections | Refer to functional diagrams to restore connections. |
| All events that may cause error 031 (problems with potentiometer and related circuits) | See error 031. |

## Alarm 034

Code 034 refers to the transmission of the serial radiographic design during movement in angling/centering. Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Inverted motor and generator | Refer to functional diagrams to restore connections. |
| Inverted dac connections | Refer to functional diagrams to restore connections. |
| All events that may cause error 031 (problems with potentiometer and related circuits) | See error 031. |

## Alarm 035

Code 035 refers to the transmission of the serial radiographic design during movement in tomography. Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Inverted motor and generator or | Refer to functional diagrams to restore connections. |
| Inverted dac connections | Refer to functional diagrams to restore connections. |
| All events that may cause error 031 (problems with potentiometer and related circuits) | See error 031. |

## Alarm 036

Code 036 refers to the transmission of the tube support during the simultaneous movement of dynamic area and angling/centering. Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Inverted motor and generator | Refer to functional diagrams to restore connections. |
| Inverted dac connections | Refer to functional diagrams to restore connections. |
| All events that may cause error 031 (problems with potentiometer and related circuits) | See error 031. |

## Alarm 037

Code 037 refers to the transmission of the serial radiographic design in particular to the motor reducer. Through the position transducer (potentiometer and related circuits) the software determines the movement without command of the transmission.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Motor reducer reversibility | Replace motor reducer. |
| All events that may cause error 031 (problems with potentiometer and related circuits) | See error 031. |

## Alarm 041

Code 041 refers to the transmission of the focal point and in particular to the position transducer (reference photocell).

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Microswitch 3FC5 not functional | Replace microswitch. |
| Power amplification board 3S3 not functional | Replace power amplification board. |
| Tie cable disconnected between photocell 3FC5 and board 3S3 | Reconnect cable. |
| Interruption in tie cable between photocell 3FC5 and board 3S3 | Replace cable. |
| Tie cable disconnected between board 3S3 and 24S9 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Interruption in tie cable between board 3S3 and 24S9 | Replace cable. |
| Incorrect connections in cables | Refer to functional diagrams to restore connections. |
| Tracks interruptions on boards | Restore connections or replace boards. |
| Rosin joints on connectors | Restore joints or replace boards. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |

## Alarm 042

Code 042 refers to the transmission of the focal point and in particular to activation 24S7.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Activation not powered or powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check tracks on board 24S9, fuses and operation of power supply board 24S5, the power supplies from the power transformer, the connections, fuses and power supplies of the transformer, line voltage. Replace board 24S9, 24S7, fuses or board 24S5, restore connections, replace fuses or power transformer 22TR1. |
| Activation powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check line voltage, the operation of power supply board 24S5, the power supplies of the transformer. Replace board 24S5 and power transformer 22TR1. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Activation blocked by thermal protection LED TER: ON LED FAU: ON | Ventilators off or not functional, use of continuous transmission function. To restore operation switch activation off, let it cool down and switch the equipment back on. |
| Motor short-circuit LED HV: ON LED FAU: ON | Check tracks of board 24S9, the tie cable of activation/motor and of motor 3MT5. Replace cable and motor. |
| Incorrect motor connection LED HV: ON LED FAU: ON | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 24S9 Input 25S2-IN4 off | Restore connections or replace board. |
| Interruption in flat-cable 24S9CFC2-25S2CN2 Input 25S2-IN4 off | Replace flat-cable. |
| Rosin joints on the connectors of board 24S9 and 25S2 Input 25S2-IN4 off | Restore joints or replace boards. |
| Input 25S2-IN4 not functional | Replace axes board 25S2. |
| Power supply failure +12Vcc for input/output board 25S2 | Check connections, power supplies on board 24S9 and 24S5, fuses on power supply board 24S5, the power supplies from the service transformer, fuses on transformer and power supplies. Replace boards 25S2, 24S9 and 24S5, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 051

Code 051 refers to the transmission of the parallel collimator and in particular to the position transducer (potentiometer and related circuits).

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Potentiometer 3PT7 not functional | Replace potentiometer. Reconnect, reacquire and check control stops following setup procedure |
| Potentiometer 3PT7 exceeds acquisition limits | Bring transmission within control stops defined limits, following setup procedure or manually. Recheck control stops. |
| Tie cable disconnected between board 25S3 and 3PT7 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |
| Interruption in tie cable between board 25S3 and 3PT7 | Replace cable. |
| Incorrect connections in tie cable between board 25S3 and 3PT7 | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 25S3 | Restore connections or replace board. |
| Rosin joints on the connectors of board 25S3 | Restore joints or replace board. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |
| Fault on reading circuit ADC7 of board 25S1 CPU | Replace board. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Power supply failure +/-15V reading circuit ADC7 of board 25S1 | Check fuses on board 25S1, the power supplies from the service transformer, the connections, fuses and power supplies of the transformer. Replace fuses or board 25S1, restore connections, replace fuses or auxiliaries transformer 22TR1. |
| Sporadic errors | Error caused by external interferences. To restore activation switch the equipment off and on. |

## Alarm 052

Code 052 refers to the transmission of the parallel collimator during movement.

Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Tie cable disconnected between board 23S8 and collimator | Reconnect cable. |
| Interruption in tie cable between board 23S8 and collimator | Replace cable. |
| Incorrect connections in tie cable between board 23S8 and collimator | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 23S10 Input 25S3-IN17 off Output 25S3-OUT23 off Dac VAN5 out of range | Restore connections or replace board. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Fault on board 23S8 | Restore and replace board. |
| Interruption in flat-cable 23S10CFC1-25S3CFC2 Input 25S3-IN17 off Output 25S3-OUT23 off Dac VAN5 out of range | Replace flat-cable. |
| Rosin joints on the connectors of board 23S9 and 25S3 Input 25S3-IN17 off Output 25S3-OUT23 off Dac VAN5 out of range | Restore joints or replace boards. |
| Input 25S3-IN17 not functional Output 25S3-OUT23 not functional Dac VAN5 out of range | Refer to testing manual to check with Test SW. Replace board I/O 25S3, CPU 25S1 and flat-cable. |
| Power supply failure +24Vcc for input/output board 25S3 | Check connections, power supplies on boards 25S3, the power supplies from the service transformer, fuses on transformer and power supplies. Replace board 25S3, restore connections, replace fuses or auxiliaries transformer. |
| Power supply failure +5Vcc or +/-12Vcc for input/output board 25S1 | Check connections, power supplies and fuses on boards 25S1, the power supplies from the service transformer, fuses and power supplies on transformer. Replace board 25S1, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 061

Code 061 refers to the transmission of the perpendicular collimator and in particular to the position transducer (potentiometer and related circuits).

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Potentiometer 3PT8 not functional | Replace potentiometer. Reconnect, reacquire and check control stops following setup procedure |
| Potentiometer 3PT8 exceeds acquisition limits | Bring transmission within control stops defined limits, following setup procedure or manually. Recheck control stops. |
| Tie cable disconnected between board 25S3 and 3PT8 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |
| Interruption in tie cable between board 25S3 and 3PT8 | Replace cable. |
| Incorrect connections in tie cable between board 25S3 and 3PT8 | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 25S3 | Restore connections or replace board. |
| Rosin joints on the connectors of board 25S3 | Restore joints or replace board. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |
| Fault on reading circuit ADC8 of board 25S1 CPU | Replace board. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Power supply failure +/-15V reading circuit ADC8 of board 25S1 | Check fuses on board 25S1, the power supplies from the service transformer, the connections, fuses and power supplies of the transformer. Replace fuses or board 25S1, restore connections, replace fuses or auxiliaries transformer 22TR1. |
| Sporadic errors | Error caused by external interferences. To restore activation switch the equipment off and on. |

## Alarm 062

Code 062 refers to the transmission of the perpendicular collimator during movement.

Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Tie cable disconnected between board 23S8 and collimator | Reconnect cable. |
| Interruption in tie cable between board 23S8 and collimator | Replace cable. |
| Incorrect connections in tie cable between board 23S8 and collimator | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 23S10 Input 25S3-IN17 off Output 25S3-OUT23 off Dac VAN6 out of range | Restore connections or replace board. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Fault on board 23S8 | Restore or replace board. |
| Interruption in flat-cable 23S10CFC1-25S3CFC2 Input 25S3-IN17 off Output 25S3-OUT23 off Dac VAN6 out of range | Replace flat-cable. |
| Rosin joints on the connectors of board 23S9 and 25S3 Input 25S3-IN17 off Output 25S3-OUT23 off Dac VAN6 out of range | Restore joints or replace boards. |
| Input 25S3-IN17 not functional Output 25S3-OUT23 not functional Dac VAN6 out of range | Refer to testing manual to check with Test SW. Replace board I/O 25S3, CPU 25S1 and flat-cable. |
| Power supply failure +24Vcc for input/output board 25S3 | Check connections, power supplies on boards 25S3, the power supplies from the service transformer, fuses on transformer and power supplies. Replace board 25S3, restore connections, replace fuses or auxiliaries transformer. |
| Power supply failure +5Vcc or +/-12Vcc for input/output board 25S1 | Check connections, power supplies and fuses on boards 25S1, the power supplies from the service transformer, fuses and power supplies on transformer. Replace board 25S1, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 071

Code 071 refers to the transmission of the iris collimator and in particular to the position transducer (potentiometer and related circuits).

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Potentiometer 3PT9 not functional | Replace potentiometer. Reconnect, reacquire and check control stops following setup procedure. |
| Potentiometer 3PT9 exceeds acquisition limits | Bring transmission within control stops defined limits, following setup procedure or manually and recheck control stops. |
| Tie cable disconnected between board 25S3 and 3PT9 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |
| Interruption in tie cable between board 25S3 and 3PT9 | Replace cable. |
| Incorrect connections in tie cable between board 25S3 and 3PT9 | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 25S3 | Restore connections or replace board. |
| Rosin joints on the connectors of board 25S3 | Restore joints or replace board. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |
| Fault on reading circuit ADC9 of board 25S1 CPU | Replace board. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Power supply failure +/-15V reading circuit ADC9 of board 25S1 | Check fuses on board 25S1, the power supplies from the service transformer, the connections, fuses and power supplies of the transformer. Replace fuses or board 25S1, restore connections, replace fuses or auxiliaries transformer 22TR2. |
| Sporadic errors | Error caused by external interferences. To restore activation switch the equipment off and on. |

## Alarm 072

Code 072 refers to the transmission of the iris collimator during movement.

Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Tie cable disconnected between board 23S8 and collimator | Reconnect cable. |
| Interruption in tie cable between board 23S8 and collimator | Replace cable. |
| Incorrect connections in tie cable between board 23S8 and collimator | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 23S10 Input 25S3-IN17 off Output 25S3-OUT23 off Dac VAN7 out of range | Restore connections or replace board. |
| Fault on board 23S8 | Restore and replace board. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Interruption in flat-cable 23S10CFC1-25S3CFC2 Input 25S3-IN17 off Output 25S3-OUT23 off Dac VAN7 out of range | Replace flat-cable. |
| Rosin joints on the connectors of board 23S9 and 25S3 Input 25S3-IN17 off Output 25S3-OUT23 off Dac VAN7 out of range | Restore joints or replace boards. |
| Input 25S3-IN17 not functional Output 25S3-OUT23 not functional Dac VAN7 out of range | Refer to testing manual to check with Test SW. Replace board I/O 25S3, CPU 25S1 and flat-cable. |
| Power supply failure +24Vcc for input/output board 25S3 | Check connections, power supplies on boards 25S3, the power supplies from the service transformer, fuses and power supplies on transformer. Replace boards 25S3, restore connections, replace fuses or auxiliaries transformer. |
| Power supply failure +5Vcc or +/-12Vcc for input/output board 25S1 | Check connections, power supplies and fuses on boards 25S1, the power supplies from the service transformer, the fuses and power supplies of the transformer. Replace board 25S1, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 081

Code 081 refers to the transmission of the parallel diaphragms and in particular to the position transducer (reference photocell).

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Photocell 3FC16 not functional | Replace photocell. |
| Power amplification board 3S3 not functional | Replace power amplification board. |
| Tie cable disconnected between photocell 3FC16 and board 3S3 | Reconnect cable. |
| Interruption in tie cable between photocell 3FC16 and board 3S3 | Replace cable. |
| Tie cable disconnected between board 3S3 and 24S9 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |
| Interruption in tie cable between board 3S3 and 24S9 | Replace cable. |
| Incorrect connections in cables | Refer to functional diagrams to restore connections. |
| Tracks interruptions on boards | Restore connections or replace boards. |
| Rosin joints on connectors | Restore joints or replace boards. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |

## Alarm 082

Code 082 refers to the transmission of the parallel diaphragms and in particular to activation 24S2.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Activation not powered or powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check tracks on board 24S9, the operation of power supply board 24S1, fuses on board 24S1, the power supplies from the power transformer, connections to transformer, fuses on transformer, the power supplies of the transformer and line voltage. Replace board 24S9, 24S2, fuses or board 24S1, restore connections, replace fuses or power transformer 22TR1. |
| Activation powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check line voltage, the operation of power supply board 24S1, the power supplies from the power transformer, the power supplies of the transformer. Replace board 24S1 and power transformer 22TR1. |
| Activation blocked by thermal protection LED TER: ON LED FAU: ON | Ventilators off or not functional, use of continuous transmission function. To restore operation switch activation off, let it cool down and switch the equipment back on. |
| Motor short-circuit LED HV: ON LED FAU: ON | Check tracks on board 24S9, the activation/motor tie cable and motor 3MT13. Replace cable and motor. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Error in motor connection LED HV: ON LED FAU: ON | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 24S9 Input 25S2-IN1 off | Restore connections or replace board. |
| Interruption in flat-cable 24S9CFC2-25S2CN2 Input 25S2-IN1 off | Replace flat-cable. |
| Rosin joints on the connectors of board 24S9 and 25S2 Input 25S2-IN1 off | Restore joints or replace boards. |
| Input 25S2-IN1 off | Replace axes board 25S2. |
| Power supply failure +12Vcc for input/output board 25S2 | Check connections, power supplies on board 24S9 and 24S5, fuses on power supply board 24S5, the power supplies from the service transformer, fuses on transformer and power supplies. Replace board 25S2, 24S9 and 24S5, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 091

Code 091 refers to the transmission of the cross diaphragms and in particular to the position transducer (reference photocell).

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Photocell 3FC17 not functional | Replace photocell. |
| Power amplification board 3S3 not functional | Replace power amplification board. |
| Tie cable disconnected between photocell 3FC17 and board 3S3 | Reconnect cable. |
| Interruption in tie cable between photocell 3FC17 and board 3S3 | Replace cable. |
| Tie cable disconnected between board 3S3 and 24S9 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |
| Interruption in tie cable between board 3S3 and 24S9 | Replace cable. |
| Incorrect connections in cables | Refer to functional diagrams to restore connections. |
| Tracks interruptions on boards | Restore connections or replace board. |
| Rosin joints on connectors | Restore joints or replace boards. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |

## Alarm 092

Code 092 refers to the transmission of the cross diaphragms and in particular to activation 24S3.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Activation not powered or powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check tracks on board 24S9, the operation of power supply board 24S1, fuses on board 24S1, the power supplies from the power transformer, connections to transformer, fuses on transformer, the power supplies of the transformer and line voltage. Replace board 24S9, 24S3, fuses or board 24S1, restore connections, replace fuses or power transformer 22TR1. |
| Activation powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check line voltage, the operation of power supply board 24S1, the power supplies from the power transformer, the power supplies of the transformer. Replace board 24S1 and power transformer 22TR1. |
| Activation blocked by thermal protection LED TER: ON LED FAU: ON | Ventilators off or not functional, use of continuous transmission function. To restore operation switch activation off, let it cool down and switch the equipment back on. |
| Motor short-circuit LED HV: ON LED FAU: ON | Check tracks on board 24S9, the activation/motor tie cable and motor 3MT14. Replace cable and motor. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Error in motor connection LED HV: ON LED FAU: ON | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 24S9 Input 25S2-IN2 off | Restore connections or replace board. |
| Interruption in flat-cable 24S9CFC2-25S2CN2 Input 25S2-IN2 off | Replace flat-cable. |
| Rosin joints on the connectors of board 24S9 and 25S2 Input 25S2-IN2 off | Restore joints or replace boards. |
| Input 25S2-IN2 off | Replace axes board 25S2. |
| Power supply failure +12Vcc for input/output board 25S2 | Check connections, power supplies on boards 24S9 and 24S5, fuses on power supply board 24S5, the power supplies from the service transformer, fuses on transformer and power supplies. Replace boards 25S2, 24S9 and 24S5, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 096

Code 096 refers to the battery of board CPU 25S1. The activation of this alarm interrupts the ignition of the equipment. All movements are disabled, the angle of tilt display flashes and shows the code while the remaining displays are switched off. This error must be reset with keys P57 or P48 to complete the equipment ignition (pulse selection key or 0° tilt). If the battery is flat the new alarms are not stored and the previous ones are lost.

| Events that may cause the error | Problem solution |
| --- | --- |
| Battery flat | Replace battery on board 25S1. |
| Fault on reading circuit of battery voltage | Replace board CPU 25S1. |

## Alarm 097

Code 097 refers to the keys, joysticks and control levers of console 1S1 and of remote keypad 3S1. The activation of this alarm interrupts the ignition of the equipment. All movements are disabled, the angle of tilt display flashes and shows the code while the remaining displays are switched off. This error must be reset with keys P57 or P48 to complete the equipment ignition (pulse selection key or 0° tilt). While the ignition is on, the movements activated by controls (pressed at ignition stage) are disabled.

| Events that may cause the error | Problem solution |
| --- | --- |
| Fault on control lever or joystick | Replace control lever or joystick. |
| Fault on connections between control levers or joysticks and console board 1S1 | Replace or restore connections. |
| Fault on console board 1S1 | Replace board. |
| Tie cable fault between console board 1S1 and board CPU console 1S2. | Replace cable. |
| Fault on CPU Console 1S2 | Replace board. |
| Fault on remote keypad board 3S1 | Replace board. |
| Fault on tie cable between control keypad board 3S1 and board I/O 25S3. | Replace cable. |
| Fault on board I/O 25S3 | Replace board |
| Tie cable fault between board I/O 25S3 and board CPU 25S1. | Replace cable. |
| Fault on board CPU 25S1 | Replace board |

## Alarm 098

Code 098 refers to the transmission of data between board CPU 25S1 and console board 1S2. The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Fault on board CPU 25S1 | Replace board. |
| Fault on Console board 1S2 | Replace board. |
| Tie cable fault between board CPU 25S1 and console 1S1 (fiber-optic cable) | Replace cable. |
| Fiber-optic cable disconnected | Connect cable. |
| Connection error on fiber-optic cable | Adjust connection. |

## Alarm 099

Code 099 refers to the absence of the command of the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Red stop button pressed on console | Restore stop button. |
| Console stop button disconnected from board 1S1 | Restore connections. |
| Tracks interruptions or rosin joints on board 1S1 | Restore connections or replace board. |
| Tie cable disconnected or incorrect connections between board 1S1 and 25S3 | Connect cable or restore connections |

| Events that may cause the error | Problem solution |
| --- | --- |
| Red stop button pressed on remote keypad | Restore stop button. |
| Tie cable disconnected or incorrect connection between IB safety, stop button and board 25S3 | Connect cable or restore connections. |
| Tilt total safety activated | Release tilt safety manually or using setup procedure. Check software and hardware control stops. |
| Tie cable disconnected or incorrect connection between tilt safety and board 25S3 | Connect cable or restore connections. |
| Elevation total safety activated (if present) | Release elevation safety manually or following setup procedure. Check software and hardware control stops. |
| Tie cable disconnected or incorrect connection between tilt safety, elevation and board 25S3 | Connect cable or restore connections. |
| IB safety activated | Release IB safety manually or following setup procedure. Check IB configuration and safety positioning. |
| Ceiling safety activated (if present) | Release ceiling safety manually or following setup procedure. Check room configuration and safety positioning. |
| Tie cable disconnected or incorrect connection between ceiling safety and board 25S3 | Connect cable or restore connections. |
| Switch 25S3SW1-4 OFF and ceiling safety absent | Turn ON switch 25S3SW1-4. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Watch-dog activated on board CPU 25S1 | Check connections of watch-dog contact and board CPU 25S1. Adjust connections or replace board. |
| Incorrect connections between board 25S1 and 25S3 (watch-dog contact) | Restore connections. |
| Tracks interruptions or rosin joints on board 25S3 | Restore connections or replace board. |
| Power supply failure +24Vcc for safety circuit on board 25S3 | Check connections, power supplies on boards 25S3, the power supplies from the service transformer, fuses on transformer and power supplies. Replace boards 25S3, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 101

Code 101 refers to the transmission of the parallel jaws and in particular to the position transducer (potentiometer and related circuits).

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Potentiometer 3PT10 not functional | Replace potentiometer. Reconnect, reacquire and check control stops following setup procedure. |
| Potentiometer 3PT10 exceeds acquisition limits | Bring transmission within defined control stops limits following setup procedure or manually and recheck control stops. |
| Tie cable disconnected between board 25S3 and 3PT10 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Interruption in tie cable between board 25S3 and 3PT10 | Replace cable. |
| Incorrect connections in tie cable between board 25S3 and 3PT10 | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 25S3 | Restore connections or replace board. |
| Rosin joints on the connectors of board 25S3 | Restore joints or replace board. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |
| Fault on reading circuit ADC10 of board 25S1 CPU | Replace board. |
| Power supply failure +/-15V of reading circuit ADC10 of board 25S1 | Check fuses on board 25S1, the power supplies from the service transformer, the connections, fuses and power supplies of the transformer. Replace fuses or board 25S1, restore connections, replace fuses or auxiliaries transformer 22TR2. |
| Sporadic errors | Error caused by external interferences. To restore activation switch the equipment off and on. |

## Alarm 111

Code 111 refers to the transmission of the perpendicular jaws and in particular to the position transducer (potentiometer and related circuits).

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Potentiometer 3PT11 not functional | Replace potentiometer. Reconnect, reacquire and check control stops following setup procedure. |
| Potentiometer 3PT11 exceeds acquisition limits | Bring transmission within control stops defined limits following setup procedure or manually and recheck control stops. |
| Tie cable disconnected between board 25S3 and 3PT11 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |
| Interruption in tie cable between board 25S3 and 3PT11 | Replace cable. |
| Incorrect connections in tie cable between board 25S3 and 3PT11 | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 25S3 | Restore connections or replace board. |
| Rosin joints on the connectors of board 25S3 | Restore joints or replace board. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |
| Fault on reading circuit ADC11 of board 25S1 CPU | Replace board. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Power supply failure +/-15V reading circuit ADC11 of board 25S1 | Check fuses on board 25S1, the power supplies from the service transformer, the connections, fuses and power supplies of the transformer. Replace fuses or board 25S1, restore connections, replace fuses or auxiliaries transformer 22TR2. |
| Sporadic errors | Error caused by external interferences. To restore activation switch the equipment off and on. |

## Alarm 121

Code 121 refers to the transmission of the longitudinal drawer and in particular to the position transducer (reference photocell).

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Photocell 3FC14 not functional | Replace photocell. |
| Power amplification board 3S3 not functional | Replace power amplification board. |
| Tie cable disconnected between photocell 3FC14 and board 3S3 | Reconnect cable. |
| Interruption in tie cable between photocell 3FC14 and board 3S3 | Replace cable. |
| Tie cable disconnected between board 3S3 and 24S9 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Interruption in tie cable between board 3S3 and 24S9 | Replace cable. |
| Incorrect connections in cables | Refer to functional diagrams to restore connections. |
| Tracks interruptions on boards | Restore connections or replace board. |
| Rosin joints on connectors | Restore joints or replace boards. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cables. |

## Alarm 122

Code 122 refers to the transmission of the longitudinal drawer and in particular to activation 24S6.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Activation not powered or powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check tracks on board 24S9, the operation of power supply board 24S5, fuses on board 24S5, the power supplies from the power transformer, connections to transformer, fuses on transformer, the power supplies of the transformer and line voltage. Replace board 24S9, 24S6, fuses or board 24S5, restore connections, replace fuses or power transformer 22TR1. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Activation powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check line voltage, the operation of power supply board 24S5, the power supplies from the power transformer, the power supplies of the transformer. Replace board 24S5 and power transformer 22TR1. |
| Activation blocked by thermal protection LED TER: ON LED FAU: ON | Ventilators off or not functional, use of continuous transmission function. To restore operation switch activation off, let it cool down and switch the equipment back on. |
| Motor short-circuit LED HV: ON LED FAU: ON | Check tracks on board 24S9, the activation/motor tie cable and motor 3MT11. Replace cable and motor. |
| Error in motor connection LED HV: ON LED FAU: ON | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 24S9 Input 25S2-IN3 off | Restore connections or replace board. |
| Interruption in flat-cable 24S9CFC2-25S2CN2 Input 25S2-IN3 off | Replace flat-cable. |
| Rosin joints on the connectors of board 24S9 and 25S2 Input 25S2-IN3 off | Restore joints or replace boards. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Input 25S2-IN3 not functional | Replace axes board 25S2. |
| Power supply failure +12Vcc for input/output board 25S2 | Check connections, power supplies on boards 24S9 and 24S5, power supplies on power supply board 24S5, the power supplies from the service transformer, fuses on transformer and power supplies. Replace boards 25S2, 24S9 and 24S5, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 131

Code 131 refers to the transmission of the transverse drawer and in particular to the position transducer (reference photocell).

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Photocell 3FC15 not functional | Replace photocell. |
| Power amplification board 3S3 not functional | Replace power amplification board. |
| Tie cable disconnected between photocell 3FC15 and board 3S3 | Reconnect cable. |
| Interruption in tie cable between photocell 3FC15 and board 3S3 | Replace cable. |
| Tie cable disconnected between board 3S3 and 24S9 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Interruption in tie cable between board 3S3 and 24S9 | Replace cable. |
| Incorrect connections in cables | Refer to functional diagrams to restore connections. |
| Tracks interruptions on boards | Restore connections or replace boards. |
| Rosin joints on connectors | Restore joints or replace boards. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |

## Alarm 132

Code 132 refers to the transmission of the transverse drawer and in particular to activation 24S8.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Activation not powered or powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check tracks on board 24S9, the operation of power supply board 24S5, fuses on board 24S5, the power supplies from the power transformer, connections to transformer, fuses on transformer, the power supplies of the transformer and line voltage. Replace board 24S9, 24S8, fuses or board 24S5, restore connections, replace fuses or power transformer 22TR1. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Activation powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check line voltage, the operation of power supply board 24S5, the power supplies from the power transformer, the power supplies of the transformer. Replace board 24S5 and power transformer 22TR1. |
| Activation blocked by thermal protection LED TER: ON LED FAU: ON | Ventilators off or not functional, use of continuous transmission function. To restore operation switch activation off, let it cool down and switch the equipment back on. |
| Motor short-circuit LED HV: ON LED FAU: ON | Check tracks of board 24S9, the activation/motor tie cable and motor 3MT12. Replace cable and motor. |
| Error in motor connection LED HV: ON LED FAU: ON | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 24S9 Input 25S2-IN5 off | Restore connections or replace board. |
| Interruption in flat-cable 24S9CFC2-25S2CN2 Input 25S2-IN5 off | Replace flat-cable. |
| Rosin joints on the connectors of board 24S9 and 25S2 Input 25S2-IN5 off | Restore joints or replace boards. |
| Input 25S2-IN5 not functional | Replace axes board 25S2. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Power supply failure +12Vcc for input/output board 25S2 | Check connections, power supplies on boards 24S9, 24S5, fuses on power supply board 24S5, the power supplies from the service transformer, fuses on transformer and power supplies. Replace boards 25S2, 24S9 and 24S5, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 133

Code 131 refers to the transmission of the transverse drawer and in particular to the microswitch of the drawer centre.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Microswitch 3FC21 not functional | Replace microswitch. |
| Power amplification board 3S3 not functional | Replace power amplification board. |
| Tie cable disconnected between photocell 3FC21 and board 3S3 | Reconnect cable. |
| Interruption in tie cable between photocell 3FC21 and board 3S3 | Replace cable. |
| Tie cable disconnected between board 3S3 and 24S9 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |
| Interruption in tie cable between board 3S3 and 24S9 | Replace cable. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Incorrect connections in cables | Refer to functional diagrams to restore connections. |
| Tracks interruptions on boards | Restore connections or replace boards. |
| Rosin joints on connectors | Restore joints or replace boards. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |

## Alarm 141

Code 141 refers to the transmission of the IB and in particular to the position transducer (reference photocell).

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Microswitch 3FC19 not functional | Replace microswitch. |
| Power amplification board 3S3 not functional | Replace power amplification board. |
| Tie cable disconnected between photocell 3FC19 and board 3S3 | Reconnect cable. |
| Interruption in tie cable between photocell 3FC19 and board 3S3 | Replace cable. |
| Tie cable disconnected between board 3S3 and 24S9 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |
| Interruption in tie cable between board 3S3 and 24S9 | Replace cable. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Incorrect connections in cables | Refer to functional diagrams to restore connections. |
| Tracks interruptions on boards | Restore connections or replace boards. |
| Rosin joints on connectors | Restore joints or replace boards. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |

## Alarm 142

Code 142 refers to the transmission of the IB and in particular to activation 24S7.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Activation not powered or powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check tracks on board 24S9, the operation of power supply board 24S5, fuses on board 24S5, the power supplies from the power transformer, connections to transformer, fuses on transformer, the power supplies of the transformer and line voltage. Replace board 24S9, 24S7, fuses or board 24S5, restore connections, replace fuses or power transformer 22TR1. |
| Activation powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check line voltage, the operation of power supply board 24S5, the power supplies from the power transformer, the power supplies of the transformer. Replace board 24S5 and power transformer 22TR1. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Activation blocked by thermal protection LED TER: ON LED FAU: ON | Ventilators off or not functional, use of continuous transmission function. To restore operation switch activation off, let it cool down and switch the equipment back on. |
| Motor short-circuit LED HV: ON LED FAU: ON | Check tracks on board 24S9, the activation/motor tie cable and motor 3MT16. Replace cable and motor. |
| Error in motor connection LED HV: ON LED FAU: ON | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 24S9 Input 25S2-IN4 off | Restore connections or replace board. |
| Interruption in flat-cable 24S9CFC2-25S2CN2 Input 25S2-IN4 off | Replace flat-cable. |
| Rosin joints on the connectors of board 24S9 and 25S2 Input 25S2-IN4 off | Restore joints or replace boards. |
| Input 25S2-IN4 off | Replace axes board 25S2. |
| Power supply failure +12Vcc for input/output board 25S2 | Check connections, power supplies on boards 24S9 and 24S5, fuses on power supply board 24S5, the power supplies from the service transformer, fuses on transformer and power supplies. Replace boards 25S2, 24S9 and 24S5, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 143

Code 143 refers to the transmission of the IB of the low IB microswitch.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Microswitch 3FC20 not functional | Replace microswitch. |
| Power amplification board 3S3 not functional | Replace power amplification board. |
| Tie cable disconnected between photocell 3FC15 and board 3S3 | Reconnect cable. |
| Interruption in tie cable between photocell 3FC20 and board 3S3 | Replace cable. |
| Tie cable disconnected between board 3S3 and 24S9 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |
| Interruption in tie cable between board 3S3 and 24S9 | Replace cable. |
| Incorrect connections in cables | Refer to functional diagrams to restore connections. |
| Tracks interruptions on boards | Restore connections or replace boards. |
| Rosin joints on connectors | Restore joints or replace boards. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |

## Alarm 151

Code 151 refers to the transmission of the grid and in particular to the position transducer (reference photocell).

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Photocell 3FC18 not functional | Replace photocell. |
| Power amplification board 3S3 not functional | Replace power amplification board. |
| Tie cable disconnected between photocell 3FC18 and board 3S3 | Reconnect cable. |
| Interruption in tie cable between photocell 3FC18 and board 3S3 | Replace cable. |
| Tie cable disconnected between board 3S3 and 24S9 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |
| Interruption in tie cable between board 3S3 and 24S9 | Replace cable. |
| Incorrect connections in cables | Refer to functional diagrams to restore connections. |
| Tracks interruptions on boards | Restore connections or replace boards. |
| Rosin joints on connectors | Restore joints or replace boards. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |

## Alarm 152

Code 152 refers to the transmission of the grid and in particular to activation 24S4.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Activation not powered or powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check tracks on board 24S9, the operation of power supply board 24S1, fuses on board 24S1, the power supplies from the power transformer, connections to transformer, fuses on transformer, the power supplies of the transformer and line voltage. Replace boards 24S9, 24S4, fuses or board 24S1, restore connections, replace fuses or power transformer 22TR1. |
| Activation powered by an out of range voltage value LED HV: OFF LED FAU: ON | Check line voltage, the operation of power supply board 24S1, the power supplies from the power transformer, the power supplies of the transformer. Replace board 24S1 and power transformer 22TR1. |
| Activation blocked by thermal protection LED TER: ON LED FAU: ON | Ventilators off or not functional, use of continuous transmission function. To restore operation switch activation off, let it cool down and switch the equipment back on. |
| Motor short-circuit LED HV: ON LED FAU: ON | Check tracks on board 24S9, the activation/motor tie cable and motor 3MT15. Replace cable and motor. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Error in motor connection LED HV: ON LED FAU: ON | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 24S9 Input 25S2-IN6 off | Restore connections or replace board. |
| Interruption in flat-cable 24S9CFC2-25S2CN2 Input 25S2-IN6 off | Replace flat-cable. |
| Rosin joints on the connectors of board 24S9 and 25S2 Input 25S2-IN6 off | Restore joints or replace boards. |
| Input 25S2-IN6 not functional | Replace axes board 25S2. |
| Power supply failure +12Vcc for input/output board 25S2 | Check connections, power supplies on boards 24S9 and 24S5, fuses on power supply board 24S5, the power supplies from the service transformer, fuses on transformer and power supplies. Replace boards 25S2, 24S9 and 24S5, restore connections, replace fuses or auxiliaries transformer. |

## Alarm 161

Code161 refers to the transmission of the longitudinal plan and in particular to the position transducer (potentiometer and related circuits).

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the layer display flashes and shows the alarm code while the remaining displays switch off.

| Possible alarm causes | Problem solution |
| --- | --- |
| Potentiometer 3PT17 not functional | Replace potentiometer. Reconnect, reacquire and check control stops following setup procedure. |
| Potentiometer 3PT17 exceeds acquisition limits | Bring tilt within defined limits, following setup procedure or manually. Recheck control stops. |
| Tie cable disconnected between board 25S3 and 3PT17 | Reconnect cable. CAUTION must be exercised with the screen earth connection. |
| Interruption in tie cable between board 25S3 and 3PT17 | Replace cable. |
| Incorrect connections in tie cable between board 25S3 and 3PT17 | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 25S3 | Restore connections or replace board. |
| Rosin joints on the connectors of board 25S3 | Restore joints or replace board. |
| Interruption in flat-cable 25S3CN2-25S1CN2 | Replace flat-cable. |
| Fault on reading circuit ADC6 of board 25S1 CPU | Replace board. |

| Events that may cause the error | Problem solution |
| --- | --- |
| Power supply failure +/-15V on the power supplies of reading circuit ADC6 in board 25S1 | Check fuses on board 25S1, the power supplies and connections from the service transformer, fuses and power supplies on transformer. Replace fuses or board 25S1, restore connections, replace fuses or auxiliaries transformer 22TR2. |
| Sporadic errors | Error caused by external interferences. To restore activation switch the equipment off and on. |

## Alarm 163

Code 163 refers to the transmission of the longitudinal plan during movement. Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the layer display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Inverted motor and generator | Refer to functional diagrams to restore connections. |
| Inverted dac connections | Refer to functional diagrams to restore connections. |
| All events that may cause error 161 (problems with potentiometer and related circuits) | See error 161. |

## Alarm 164

Code 164 refers to the transmission of the longitudinal plan and in particular to the motor reducer.

Through the position transducer (potentiometer and related circuits) the software determines the movement without command of the transmission.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the layer display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Motor reducer reversibility | Replace motor reducer. |
| All events that may cause error 161 (problems with potentiometer and related circuits) | See error 161. |

## Alarm 189÷196

Codes 189÷196 refer to board CPU 25S1.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Board CPU 25S1 not functional | Replace board CPU. |

## Alarm 197

Code 197 refers to the axes board for stepper motors 25S2.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Axes board 25S2 not functional | Replace axes board. |

## Alarm 198

Code 198 refers to the transmission of data between board CPU 25S1 and PC in setup mode in <Telediagnosis>. The activation of this alarm disables the power transformer connection relay. All movements are disabled, the layer display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Serial cable or connection by modem interrupted | Recheck cable. To restore operation switch the equipment off and on. |

## Alarm 199

Code 199 refers to the transmission of data between board CPU 25S1 and PC in telecontrol mode in <Telediagnosis>. The activation of this alarm disables the power transformer connection relay. All movements are disabled, the layer display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Serial cable or connection by modem interrupted | Recheck cable. To restore operation switch the equipment off and on. |

## Alarm 200

Code 200 refers to the fault of Eeprom mounted on board 25S1.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Eeprom not functional | Replace Eeprom. Reinsert default data on Eeprom in setup mode. Acquire configuration data, control stops and calibration following the service manual. |
| Eeprom absent | Insert Eeprom. |

## Alarm 201

Code 201 refers to the content of Eeprom mounted on board 25S1.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Eeprom new (during testing) | Reinsert default data on Eeprom in setup mode. Acquire configuration data, control stops and calibration following the service manual. |
| Replace board CPU 25S1 | Remove Eeprom from the replaced board. |
| Eeprom content modified due to change in software version | Follow instructions enclosed with software version. |
| Eeprom dirty | Reinsert default data on Eeprom in setup mode. Acquire configuration data, control stops and calibration following the service manual. |

## Alarm 202

Code 202 identifies the activation of a hypothetical alarm that has been generated but not dealt with by the application software. There are no plausible events that may activate this alarm.

The activation of this alarm disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.