Table 2 - Voltage Tolerance Specifications 

Voltage  Low Voltage  High Voltage 

3.65 3.61  3.68

5.35 5.24  5.45

-5.35 -5.45  -5.24

15.00 14.70  15.30

-15.00 -15.30  -14.70

48.0 47.0  49.0

12.00 11.80  12.20

 Step 3.1 – Compare each of the initial voltage 

measurements to the low voltage and high voltage 

values shown in Table 2. 

A voltage adjustment is necessary: 

•  If the initial voltage measurement is less than the 

low voltage specification shown in Table 2. 

•  If the initial voltage measurement is greater than 

the high voltage specification shown in Table 2. 

•  If the power supply module is actively drifting 

(even if it is within the voltage tolerances 

shown in Table 2). 

If the power supply is not drifting and the voltages 

are within the low and high limits shown in  

Table 2 then do not adjust the power supply.  

Note: the voltage tolerance numbers used in Table 2 

come from the Power Supply Module 

Specifications. 

